Revision of Project Success Criterion

Change Control Record — supersedes the success metric stated in the original project briefing.

| Document type | Change Control Record / Scope Clarification |
| --- | --- |
| Project | Insole Foot Classification System |
| Subject | Revision of the model’s original Success Criterion |
| Authorized by | Mohammed Fachry Dwi Handoko |
| Status | PENDING |
| Supersedes | “Classification (accuracy >90% for unseen images)” in original brief |

# 1. Purpose

This record formally revises the success criterion for the AI Insole Foot Classification System. The original brief required image-based classification with accuracy above 90% on unseen images. Empirical testing during development demonstrated that this criterion is not achievable from the supplied data, for a structural reason described below.

This document states the revised criterion, the evidence supporting the change, and the resulting system design. It is intended to be retained in the project file as the authoritative reference for what the delivered system does and does not claim.

# 2. Original Criterion

The original project brief specified, under EXPECTED OUTPUT:

“Classification (accuracy > 90% for unseen images).”

This criterion assumed that foot arch class is recoverable from the supplied photographs. Development evidence shows this assumption does not hold for this dataset.

# 3. Evidence

The five clinical classes are defined by the project brief in terms of arch height in centimetres (Severe Flat < 2.7; Flat 2.7–3.5; Normal 3.6–5.5; High 5.6–6.4; Severe High > 6.4). The ground-truth label is therefore a deterministic function of a single measured quantity. A controlled diagnostic (retained in the codebase as scripts/diagnose_model.py) evaluated the trained model on a held-out test split under three input conditions:

| Evaluation condition | Accuracy | Verdict |
| --- | --- | --- |
| Rule applied to TRUE measurements (no ML) | ~100% | Exact |
| Model with measurements present | ~88% | Measurement-driven |
| Model images-only (measurements withheld) | ~33% | Near chance |

The image-only condition collapsed to approximately chance level, with 0% recall on the Normal and High classes. The supplied photographic views (top/AP, posterior/heel, and lateral as provided) do not reliably encode arch height. A model asked to predict arch class from these images cannot succeed, and earlier apparent successes were attributable to dataset artefacts rather than learned anatomy.

# 4. Revised Success Criterion

The system shall classify foot type deterministically and exactly (≥ 99%) from the clinical arch-height measurement, applying the project’s defined centimeter bands.

Where a measurement is not supplied, the system shall provide an image-based estimate, clearly and unmistakably flagged as assistive and non-authoritative, and shall not present it as a measured result.

Because the classes are defined by measurement thresholds, a deterministic rule on the measurement is not merely adequate - it is exact by construction, explainable, and clinically defensible. This is a stronger guarantee than a probabilistic image model could ever provide, and it matches how the ground-truth labels were themselves produced.

# 5. System Behavior

- Measured path (authoritative): when the clinician supplies arch height, classification is deterministic, exact, and shown with a green “MEASURED” banner at 100% confidence.

- Estimated path (assistive): when no measurement is supplied, arch height is estimated from images, the result carries an amber “ESTIMATED — NOT a clinical measurement” banner, confidence honestly reflects uncertainty, and the report instructs the user to confirm with a real measurement.

- Provenance is always explicit: every result records and displays which path produced it, so a clinician is never misled about trust level.

- Retained ML value: the neural network continues to provide image-based measurement estimation (as an assistive pre-fill) and the generative insole-configuration output.

# 6. Impact

This revision changes the stated metric, not the deliverable’s clinical usefulness. The delivered system is more reliable than the original criterion implied: classification is exact when measurements are available (the normal clinical workflow), and honestly bounded when they are not. No downstream component depends on the superseded 90%-image metric.

# 7. Authorization

As the Project Owner, the undersigned individual approves the revised success criterion in Section 4 as superseding the success metric in the original project brief and accepts the system behavior described in Section 5 as the basis for delivery and acceptance.

| Name | Signature | Date |
| --- | --- | --- |
| Mohammed Fachry Dwi Handoko |  | 05 / 17 / 2026 |
| Role: Project Owner & Engineer - Insole Foot Classification System |  |  |

Supporting evidence is reproducible at any time by running scripts/diagnose_model.py and scripts/verify_dataset.py against the delivered system. This record should be read together with those outputs.