Verification & Test Report

Insole Foot Classification System - evidence record for delivery and acceptance.

| Document type | Verification & Test Report |
| --- | --- |
| Project | Insole Foot Classification System |
| Version | 1.0 |
| Status | COMPLETED |
| Related documents | Project Brief; Revision of Success Criterion; SRS |
| Reproducibility | All results are regenerable via scripts/verify_dataset.py and scripts/diagnose_model.py |

# Contents

Contents	1

1. Summary	2

2. Test A — Measured Path (authoritative)	2

3. Test B — Estimated Path (assistive)	3

4. Diagnostic Probes	3

5. Dataset Verification	4

6. Known Limitations (verified, documented)	4

7. Authorization	4

# 1. Summary

This report records the verification evidence for the delivered system. The core classification function was validated under both operating paths on a known patient, and the underlying data and model behavior were characterized by two reproducible diagnostic scripts. The evidence supports acceptance against the revised success criterion: deterministic, exact classification from clinical measurements, with a flagged assistive image path.

Overall verdict: PASS

The measured path is exact by construction and verified; the estimated path behaves honestly and is correctly labelled non-authoritative; known limitations are documented and reproducible.

# 2. Test A — Measured Path (authoritative)

Objective: confirm that - when a clinical arch-height measurement is supplied - classification is deterministic and matches the project’s defined bands.

Subject: Patient P014

Ground truth from the consolidated measurement sheet: calcaneal inclination 24.6°, heel angle 2.6°, arch height 4.69 cm, kite angle 31.5°, 1st metatarsal–talus 3.8°.

True label: Normal Foot (4.69 cm falls in the Normal band 3.6–5.5 cm).

| Inputs | Three views (lateral, AP, heel) plus all five measurements entered |
| --- | --- |
| Expected | Normal Foot, 100% confidence, authoritative indicator |
| Observed | Normal Foot, 100.0% confidence; green “MEASURED — authoritative” banner; all five values tagged (measured); other classes 0.0% |
| Result | PASS |

The interpretation here is that the measured path does not consult the network’s class output; it instead applies the brief’s bands deterministically. The result is therefore exact by construction and will hold for any patient with a supplied arch-height measurement.

# 3. Test B — Estimated Path (assistive)

The objective here is to confirm that - when no measurement is supplied - the system estimates from images, labels the result as non-authoritative, and expresses honest uncertainty.

| Inputs | Same three P014 views; all measurement fields cleared |
| --- | --- |
| Expected | Result flagged assistive; sub-100% confidence; confirmation warning |
| Observed | Normal Foot at 81.9% confidence; amber “ESTIMATED — NOT a clinical measurement” banner; estimated arch height 5.03 cm; 16.6% mass on High Arch; values tagged (est.); warning to confirm with a real measurement |
| Result | PASS |

The interpretation here is that the prediction was correct, but the system did not overstate it. The estimate (5.03 cm) deviated from truth (4.69 cm) by 0.34 cm and the residual probability mass honestly reflected proximity to the High-Arch boundary. This is the desired assistive behavior, and not an authoritative claim.

# 4. Diagnostic Probes

Source: scripts / diagnose_model.py

Code is evaluated on the held-out test splitting session. These probes characterize why the system is measurement-first.

| Probe | Accuracy | Reading |
| --- | --- | --- |
| Rule on TRUE measurements (no ML) | ~100% | Exact |
| Model with measurements present | ~88% | Measurement-driven |
| Model images-only (measurements withheld) | ~33% | Near chance |
| Measurement-rule baseline (train/val/test) | 99–100% | Ceiling |

Conclusion from the probes indicates that the arch height is not visually recoverable from the supplied photographic views (images-only collapses to ~33% with 0% recall on Normal and High). The classification task is exactly solved by a deterministic rule on the inputted measurement, which is the design that the system implements.

# 5. Dataset Verification

Source: scripts/verify_dataset.py. Confirms the training data is sound and free of patient leakage.

- 1,398 patient samples matched (image + measurement); 5 of 5 classes were populated.

- Class distribution: Severe Flat 173, Flat 325, Normal 500, High 260, Severe High 140.

- View coverage: lateral 100%, top 100%, back 100%.

- Unique patients 1,398; duplicate patient IDs 0 (no train/test leakage by patient).

# 6. Known Limitations (verified, documented)

- Image-only classification is unreliable for this dataset and is exposed only as an assistive, marked or flagged path.

- Estimated measurements are approximate and require clinical confirmation before use.

- Results depend on the project’s fixed arch-height bands; changing them requires a central configuration change.

# 7. Authorization

As the Project Owner, the undersigned accepts the verification evidence in this report as sufficient for delivery and acceptance of the AI Insole Foot Classification System.

| Name | Signature | Date |
| --- | --- | --- |
| Mohammed Fachry Dwi Handoko |  | 05 / 17 / 2026 |
| Role: Project Owner & Engineer - Insole Foot Classification System |  |  |

Every result in this report is reproducible on demand from the delivered codebase. This report should be read together with the System Requirements Specification and the Revision of Success Criterion documents, respectively.